Standard Operating Procedure (SOP), Instron Mechanical Tester

THIS INSTRUMENT RREQUIRES TRAINING BY INSTRUCTOR

SOP ver.1 (2024.08.27)

Instron Mechanical Tester has the following risks. Check Risk Assessment before starting to use!

- A hand caught in the machine

- A finger caught in the clamps

- Load cell break by operation without safety limit setting

- Cutting your hand with a cutter while preparing the samples

Starting up the Instron

- Check air supply

  - Air pressure should be more than 6 bar

- Check the load cell and clamps

  - Standard load cell: 5 kN new load cell (± 0.5% of reading down to 5 N)

  - Standard clamps: New tensile clamp, Old tensile clamp or Compression fixture

  - If you would like to use a non-standard clamp system, contact the machine responsible

- Turn on Instron (Right rear of the machine)

- Turn on PC and log in (Account: lab-user, Password: no password)

- Open Instron software from PC and log in (Account: CHALMERS, Password: CHALMERS)

New tensile clamp, Old tensile clamp, Compression fixture

Method Preparation

- Click “Methods”

- Select a method or create a method from zero.

The software is very complex, so we recommend using our Template Method.

- “Template Method, New tensile clamp”

- “Template Method, Old tensile clamp”

- “Template Method, Compression Fixture”

If you start the method from zero, the load cells may crash each other during the test or auto-positioning, depending on the setting. Contact instructors or experienced staff. Use the safety lock system and pay maximum attention when you use your method's auto-positioning function for the first time.

- Modify “Fixture Length” in “Sample” tab depend on your test condition

- The rest of the section can be modified at this time but also can be modified during the test

- Template Method already includes the clamps length information (XX tab), so don’t touch here.

- The standard “Grip air pressure” is 4.0 bar. This should not be 6.0 or more. This will lead to errors.

- Save as the method with your new name (you can keep using this next time)

Test Measurement

- Click the “Test”

- Open created your method

- Auto-position: software shows you the displacement to go to the “Fixture Length” you set in your method

- Check the displacement will not make the clamps crash each other!!

- Move safety lock system positions so as not to make the clamps crash.

- Press the “Unlock” and then the “Return” button

- Now, the clamp is already in the starting position

- Input test information

  - Sample

- Input a sample information

  - Width

  - Thickness

  - Deformation rate

- Start a test with the press of the “Unlock” and “Start” buttons on the machine

- Once you get a result, leave a memo following the pop-up instructions

- Repeat measurements

- Calculation conditions, etc., can be changed during the test using the Method tab. Make sure to save the method after the modification.

Finish the Measurement

- Click the “Finish” icon

- The raw data, PDF report (the style can be changed from method), and CSV files are automatically saved into the folder you set during the test.

- Export data to your PC through the internet (you can leave your data on this computer but not guarantee the backup)

- Close the software

- Turn off PC (only recommended if it is before a long holiday)

- Turn off Instron, Right rear of the machine (only recommended if it is before a long holiday)

NOTIFY THE INSTRUMENT RESPONSIBLE OF ANY ERRORS OR MISFUNCTIONING