## Slide 1: High Performance Lambdas
Steve Upstone UV/Vis Product Specialist
May 2019

## Slide 2: Electromagnetic Spectrum

## Slide 3: Comparison Between UV and IR
UV Absorption Spectrum
IR Transmittance Spectrum

## Slide 4: Some of PerkinElmer’s History in UV/Vis and UV/Vis/NIR
137UV (1957)
402 (1967)
57X Series (1977)
Lambda 9 with 7500 32bit computer (1984)

## Slide 5: High Performance Lambda Instruments
UV/Vis/NIR | 175-3300 | Up to 8A | 2 or 3 detector system (PMT/PbS or PMT/InGaAs/PbS)
UV/Vis | 175-900nm | Up to 8A | PMT Detector

## Slide 6: System Concept – Two Sampling Areas

## Slide 7: Snap-in Sampling Modules
Transmission
URA

## Slide 8: Inside the HP Lambda Instrument (2D Lambda 1050 shown here)

## Slide 10: Chopper System (Beam Separator)

## Slide 11: Why Is a Double Beam System Necessary?
The reference beam (usually the rear position) is necessary to compensate for changes in lamp energy (which would cause the signal to fluctuate or drift)
The double beam system is also useful when using integrating spheres as it helps to remove some systematic “substitution” errors caused by modification of the sphere behaviour by the sample.

## Slide 12: UV/Vis Detector - Photomultiplier
Detector based on a phototube which converts photons to electrons as they strike the phtocathode. These electrons are amplified by the dynode chain until they arrive at the anode.
Photomultiplier requires high voltage (typically around 900V) to operate. This is the gain
When operating in double beam mode (eg Absorbance, transmittance or reflectance), the gain is set automatically by the instrument.
When using either of the two single beam energy modes (E1, E2), the (relative) gain must be set manually via a factor (a value between 0 and 255)
Gridless Hamamatsu R6872

## Slide 13: Photomultiplier Response Curve
The HP Lambdas use the gridless R6872 photomultiplier detector which is rated down to 160nm (but the instruments are rated down to 175nm with N2 purging).
The sensitivity of both detectors drops off rapidly above 800nm

## Slide 14: Lead Sulphide (PbS) Detector
The lead sulphide (PbS) detector is used for all wavelengths above the detector change point (normally 860.8 nm)
The detector is cooled to 16°C (using a single-stage Peltier element) to give a constant temperature environment (and, therefore, constant instrument characteristics)
It is a small detector relative to the photomultiplier

## Slide 15: InGaAs Detector technology improves NIR detection
Typical InGaAs performance
Pbs detector typical response

## Slide 16: Indium Gallium Arsenide (InGaAs) Detector
InGaAs detector designed to work between 860 and 2500nm
Detector cooled to -50oC (thermoelectrically)

## Slide 17: Superior sensitivity of InGaAs compared to PbS for NIR
Pbs Detector / Fixed 2nm slit
WB InGaAs Detector / Fixed 2nm Slit

## Slide 18: InGaAs NIR Linearity

## Slide 19
The InGaAs detector provides the best all a round performance
In the range of 860-2500nm.
The majority of applications for High End Spectrophotometers
covers the region from 175nm to 2500nm

## Slide 20: Lambda performance on narrow band (1nm) filter
Pbs Detector with 0.4nm Fixed slit
Experiment w/ InGaAs 0.4nm Fixed Slit

## Slide 21: Scan Speed, Response and Data Interval
Steve Upstone

## Slide 22: Scan Speed, Response and Data Interval
How the Instrument Works
Chopper Segment Cycle Signal Correction (CSSC)
Trading Rules between Response, Data Interval and Scan Speed

## Slide 23: Chopper System Produces Sample and Reference Beams

## Slide 24: Conventional Chopper (Lambda 9 from 1984) – Scan Speed Dependent Wavelength Shifts if Grating Moves Continuously

## Slide 25: Why This Error Occurs
If the monochromator is moving continuously during measurement then the faster the scan speed, the more the grating will have moved between sample and reference measurements.
This time offset produces wavelength shifts and impacts on the accuracy of the system
If the grating is in the same position for both sample and reference readings, the error is eliminated.

## Slide 26: CSSC (Step Scan) Chopper System – Error is Eliminated

## Slide 27: CSSC* System – Read and Move Grating on Alternate Cycles
*Chopper Segment Signal Correction

## Slide 28: Chopper Cycle and Response
Each complete reading cycle takes 40 milliseconds (ie 0.04 seconds)
Data can be averaged by increasing the response time (also known as the SAT – Signal Averaging Time) so, for example, a response time of 0.16 seconds would be the average of four readings (ie 4 x 0.04)
Averaging data improves noise according to the square root law that states that the noise will be improved by the square root of the number of readings

## Slide 30: Square Root Law
Response (SAT) Time (seconds) | Number of Readings Averaged | Expected Improvement in Signal to Noise
0.04 | 1 | 1
0.16 | 4 | 2
0.64 | 16 | 4
2.56 | 64 | 8

## Slide 31: Scan Speeds for a 0.16 Response Time at Different Data Intervals
Data Interval (nm) | Expected Scan Speed (nm/min)
0.1 | 37.4
1 | 361.0
2 | 696.0
5 | 1570.0
10 | 2700.0

## Slide 32: Trading Rules
Out of
  Fast Scanning
  Low noise
  Small data interval
You can choose TWO parameters (ie if you want a small data interval and lower noise, this will result in a slower scan speed)

## Slide 33: The Detector “Step”

## Slide 34: The Detector “Step”
The detector step normally has its roots in polarisation. Detectors are sensitive to polarisation
Mainly observed
It can also occur due to sample refraction
As standard (ie unless polarisers or depolarisers are used) the instrument produces a polarised beam with polarisation being strongest at the “Wood’s Anomalies” features
Although the step is the most obvious manifestation of polarisation, it can affect the whole spectrum
It cannot be removed at the touch of a button (although its effects can be masked by a combination of software algorithms and smoothing). It is better to understand why it is happening and remove it fundamentally

## Slide 35: Inherent polarisation from the instrument (mainly from gratings but also mirrors)
Wood Anomalies

## Slide 36: Effect of Angle of Incidence on Polarisation
Samples measured at angles of incidence greater than (approx.) 15 degrees react differently to vertically and horizontally polarised light
This effect is greatest close to the Brewster Angle

## Slide 37: Common Beam Depolariser Accessory (CBD)
The Common Beam Depolariser (CBD) Accessory installs inside the optical compartment of the instrument and is switched in and out through software

## Slide 38: Detector Step and Polarisation State - Reflectance

## Slide 39: Detector Change and Polarisation State - Transmittance

## Slide 40: Detector Step and Refraction
Image on detector during autozero (no sample)
Sample introduced – refraction causes beam to illuminate different part of detector (which may not be completely homogeneous)
Active area of detector (shown in grey)

## Slide 41: Using an Integrating Sphere to Eliminate the Detector Change
In situations where refraction and/ or multiple internal reflection are causing the beam to shift or change shape, an integrating sphere can be used as a detector
It can be thought of a nearly perfectly homogeneous detector
It is also useful for use with accessories where the beam is inverted between the baseline and sample measurement (eg VN type accessories) or to reduce alignment issues with more complex optical arrangements.

## Slide 42: 4x Rule
There are four times fewer lines per mm on the NIR grating than the UV/Vis (360 lines/mm cf 1440 lines/mm)
This means that, for the same physical slit width (ie in mm), that the resolution will be 4x larger (eg a 2nm slit in the UV/Vis is equivalent to an 8nm slit in the NIR)
By keeping the physical slit width the same, the polarisation balance is more likely to be conserved

## Slide 43: General Recommendations
Polarisation must be controlled at angles of incidence greater than 15 degrees
The “4x rule” minimises changes in polarisation at the detector change by maintaining the same physical beam size
Consider using an integrating sphere as a detector for highly refractive samples

## Slide 44: Beam Attenuation for Measuring at High Absorbance
Steve Upstone
November 2009

## Slide 45: Overcoming Noise Issues at High Absorbance
Stacked Didymium Filters – no beam attenuation
Notes: This is a common problem – trying to measure at greater than 4AU without using reference beam attenuation.

## Slide 46: Using Motorised Beam Attenuators to Reduce Noise at High Absorbance
Reference Beam Attenuation (RBA) is a technique to minimise noise when measuring highly absorbing samples
Motorised sample and reference beam attenuators

## Slide 47: Full Sequence

## Slide 48: Sample without Reference Beam Attenuation

## Slide 49: 1% Reference Beam Attenuation – Manual Mode

## Slide 50: Automatic Mode Beam Attenuators
Gives the best signal to noise for high absorbance/ low transmittance measurements
Three measurements taken – software automatically calculates and removes contribution from the attenuator

## Slide 51: 0.1% “Automode” Beam Attenuation (Lambda 1050)

## Slide 52: Automatic Mode Reference Beam Attenuation for Low Noise at High Absorbance
This slide shows the superior results at high absorbance when using the unique automatic mode of reference beam attenuation

## Slide 53: Maintenance

## Slide 54: Maintenance
Calibration
Service Methods for checking performance
Changing lamps

## Slide 55: Calibration
Calibration is a procedure for checking and correcting for any small drift in slit and wavelength accuracy
If your instrument has a standard detector (either 2D or 3D) the procedure is best done using that
The procedure takes between 5 and 10 minutes

## Slide 56: Service Methods
Service methods are used to check instrument performance
There is a large “library” of tests that check things like wavelength accuracy, stray light, absorbance accuracy etc.
Some tests require test materials (e.g. glass filters or test solutions) but others do not.

## Slide 57: Sampling Options for the Lambda 1050

## Slide 58: Lambda X50 Series with 150mm Integrating Sphere

## Slide 59: Lambda X50 Series with TAMS* Goniometer
* Total Automated Measurement System

## Slide 60: General Purpose Optical Bench (GPOB)

## Slide 61: URA Accessory for Measuring Specular Reflectance of Small Components

## Slide 62: Oxford Instruments Optistat CF and DN Cryostats

## Slide 63: Questions?
Notes: Questions