Lambda 1050 – Standard Operating Procedure

- If you need to install hardware accessories, always do this with the instrument switched off.

- Switch on the instrument, ideally 1 hour before your experiment to allow time for lamps to warm up and stabilize.

- Start the UV WinLab Software. Login as Administrator, no password required.

- Select a method from the list, for example one of the “Chalmers” methods. Note: right click and select Edit mode if you want to modify the method.

- Please create your own folder under User methods, and save your methods there.

- Check your method settings:

  - Data collection

    - Set your measurement range

    - Set your ordinate mode (Absorbance, Transmittance or Reflectance)

- Do a background measurement:

  - Check that the beam paths are free from obstruction

  - Close the sample compartment and click Autozero
Or:

  - Place your reference in the sample position

  - Close the compartment and click Autozero (if you do this, keep reference position empty during your measurements)

- Sample info in method:

  - Adjust the number of samples you plan to characterize

  - Enter the sample ID and description (this can be done for all samples at once or as your experiments progress)

- Run your experiment:

  - Place your sample in the sample position and, if relevant, place a reference in the reference position.

  - Click Start.

  - If you have more than one sample: exchange your sample and click OK to use the same settings, or click cancel and modify your settings.

- Save your data:

  - Select File -> Save Results -> New Task. The task includes all the raw data and csv-files. After saving, move your task (using the UV WinLab Explorer) to your own folder under User tasks.

  - Optionally, right click on the spectra in the Graph window and select save. Move your data to the Save UVdata here folder on the Desktop.

- Finish:

  - Switch off the instrument. Remove any accessories you added.